## Slide 1
Nové možnosti rychlého a ultracitlivého panelového NGS testování mutací
ve tkáňových a tekutých biopsiích s využitím HANDLE® a ddCAP®
technologií přípravy knihoven
Marek Minárik

## Slide 2
Disclosure information
Ownership in Elphogene, s.r.o. company
Owership in Carolina Biosystems, s.r.o. company

## Slide 3
Mutation profile
A mutation informed liquid biopsy
Targeted detection
in plasma
Plasma profiling
Tissue profiling

## Slide 4
A two-tier approach
velký NGS panel
malý NGS panel

## Slide 5: NGS (panel) sequencing technologies

## Slide 6
Amplicon sequencing Hybridization capture sequencing:
HANDLE®
Halo-shape ANnealing and
Defer-Ligation Enrichment
ddCAP®
Dual-directional capture technology

## Slide 8
Clean-up
~30min
5’
mRNA
3’
5’
3’
cDNA
Reverse Trasncription
~60min
Hybridization
~2h
Extension-Ligation
~10min
Digestion
~40min
PCR
~60min
cDNA or gDNA
Simple workflow and low cost
One-Tube, Hands-on time < 1 h, total library construction time < 6 h
Easy to implement in local molecular lab
8
Extension primer
UMI 1
Linker-Tag1
Linker-Tag2
UMI 2
Ligation primer
UMI (Unique Molecular Identifier): For error correction to improve sensitivity and accuracy.
  Classic HANDLE® - Library Prep
Sensitivity: 1% MAF

## Slide 9
Step 1.
Nucleic acid extraction
Step 2.
Library preparation
Step 3.
Sequencing
Step 4.
Bioinformatics data analysis
TAT: ~5 h
TAT: ~6 h
TAT: ~30 h
TAT: ~1 h
Day 1
Day 3
Day 2~3
Day 2
9
Extraction
Sample collection
Nucleic acid extraction
Quality control
Library Prep.
Sample barcoding
Target enrichment
Amplification/purification
Quality control
Sequencing
PE150 reads
Quality control
Bioinformatics
(Automatic data analysis)
Variant calling
Variant annotation
Quality control
ThermoCycler
Magnetic Stand
Fluorometer
or
Fragment analysis
Miseq/MiseqDx
Nextseq 500/550Dx
  Classic HANDLE® workflow

## Slide 10
Parameter | Specifications
Technology | HANDLE®
Target Regions | 40 genes and MSI
Alterations Detected | SNV, Indel, Fusion, CNV, MSI
Tumor Type | Pan-cancer
Sample Type | FFPE tumor tissue
NA Requirement | DNA: minimum 50 ng (Optimal 70 ng) / RNA: minimum 30 ng (Optimal 400 ng)
Limit of Detection (LoD) | 1% allele frequency; 5% allele frequency in non-core regions; minimum 20% tumor content
Recommended Data Amount per Sample | MiSeq/MiniSeq: 375 Mb/sample NextSeq: 1 Gb/sample
Sequencing Type | PE150
Sequencer | Illumina MiSeqDx, NextSeq 550Dx
TAT for Library Preparation | 6 h (hands-on time 1 h )
TAT from Sample to Report | 3 days
10
  Classic HANDLE® NGS Panel Specs

## Slide 11
Dual: probes targeting double strands
sense strand capture
antisense strand capture
11
  Classic ddCAP® - Library Prep
Notes: 1.double UMI,reduce error

## Slide 12
ddCap® Comprehensive Panel Workflow
Step 1.
Nucleic acid extraction
Step 2.
Library preparation
Step 3.
Sequencing
Step 4.
Data analysis
Extraction
Sample collection
Nucleic acid extraction
Quality control
Library Prep.
Sample barcoding
Target enrichment
Amplification/purification
Quality control
Sequencing
PE150 reads
Quality control
Bioinformatics
(Automatic data analysis)
Variant calling
Variant annotation
Quality control
TAT: ~5 h
TAT: ~2 d
TAT: ~25-44 h
TAT: ~3-5 h
TAT: ~4 h
Day 1
Day 5
Day 3~4
Day 2~3
12
ThermoCycler
Ultrasonicator
(Covaris M220)
Magnetic Stand
Fluorometer
or
Fragment analysis
Nextseq 500
Novaseq 6000

## Slide 13
Parameter | Specifications
Technology | ddCAP®
Target Regions | 128 genes and MSI
Alterations Detected | SNV, Indel, Fusion, CNV, SNP, MSI *
Tumor Type | Cross-tumor
Sample Type | FFPE tumor tissue, liquid biopsy
DNA Input | FFPE DNA: optimal 100 ng (minimum 50 ng) Plasma cfDNA: optimal 30 ng (minimum 10 ng)
Limit of Detection (LoD) | FFPE DNA: 5% allele frequency; 20% tumor content Plasma cfDNA: 0.5% allele frequency
Data Output per Sample | FFPE DNA: 1.5 Gb/sample Plasma cfDNA: 8 Gb/sample
Sequencing Type | PE150
Sequencer | Illumina NextSeq 500, NovaSeq 6000
TAT for Library Preparation | 2 d (hands-on time 4 h )
TAT from Sample to Report | 5 days
ddCAP® Comprehensive Panel – Specifications
13
* CNV and MSI detection are for tissue sample only.

## Slide 14
14
ddCAP® Comprehensive Panel – genes

## Slide 15
ddCAP® Comprehensive panel kit (129 genů, 5 dní TAT, 24 reakcí) €337,- (Kč 7755,-)/vzorek
15
Plasma
ddCAP® Comprehensive panel kit (129 genů, 5 dní TAT, 12 reakcí) €490,- (Kč 11633,-)/vzorek
Tkáň
Classic HANDLE® NGS kit (40 genů, 3 dny TAT, 24 reakcí) €175,- (Kč 4136,-)/vzorek

## Slide 16
16